## CYT
PROBLEMA | CAUSA | SOLUZIONE
The system does not switch on | Power cable not connected. Main electrical socket not on. Circuit Breaker not on. Broken Main Fuse | Connect mains cable. Turn on main socket. Check the circuit breaker, turn on.
Low output power | Damaged optical fiber | Inspect the Fiber. Replace the Fiber.
No output power | The laser is not in ready mode. Footswitch broken. | Select Ready mode.
Alarm: footswitch not connected |  | Connect the footswitch.
Remote door interlock not connected | interlock not connected | Connect the interlock.
Fiber: warning | Fiber not connected properly. | Re-connect the fiber, tighten securely.
Alarm chillers 1 or 2 | Chiller 1 or 2 not working. Low water level. Low flow. | Replace the water filter. Add water
Other Alarms: SHUTTER FIBER CURRENT CHECK INTERLOCK 1 INTERLOCK 2
Low Energy | Energy mesured by internal Powermeter is <20% | Check current tables
Very Low Energy | Energy mesured by internal Powermeter is <30% | Check current tables

## LHT
PROBLEMA | CAUSA | SOLUZIONE
The device can not start up | • Power line not connected • Main line breaker is off • Emergency push button is pressed down | • Check line cord breaker • Pull up the emergency push button
Footswitch Error | • Footswitch disconnected • Defective Footswitch | • Connect the footswitch • Check if footswitch must be replaced
Water Flow Error | • Not enough water • Water filter occluded • Water tube twisted • Water leakage | • Fill water reservoir • Check water filter • Check water tubes • Look for any water leakage
Water Temperature Error | • Room temperature too high or air inlet/outlet not free | • Check cooling circulation and fans • Assure that a correct distance from the rear panel to the wall exists • Check water pump
Flashlamp error | • Flashlamp is exhausted or broken • Power supply | • Check power supply • Exchange the flashlamp
Simmering Error | • Simmering breakage • Flashlamp broken • Trigger transformed failed | • Check the compact power supply (see Ch. 6.4.1) • Check the flashlamp • Check trigger transformer
OverVoltage/Temp Error | • Problem in power electronics | • Check the compact power supply
External Interlock Error | • Interlock not well connected or not connected at all | • Check the connection of the Interlock
Energy Warning: high or low energy | • Fluctuation of the energy • Incorrect energy | • Check the photodidode reading • Check the optics and laser head
No correspondence between displayed energy and application effects | • Damage of the fiber | • Exchange the fiber
Shutter Error | • Incorrect shutter position | • Change the shutter or its electronic board

## CYH
PROBLEMA | CAUSA | SOLUZIONE
Key switch does nothing |  Power line is not connected Main line breaker is off Emergency red push button is pressed  Main line input/low voltages connections are not proper  Low voltages Power supply (PS) is damaged  Microprocessor or embedded PC boards are damaged |  Check main line breaker (room and machine)  Defeat the emergency push button  Check main line filter I/O voltage  Check key and emergency push button I/O voltage  Check low voltage PS I/O voltages  Check low voltages on microprocessor board and embedded PC  Substitute damaged parts
No correspondence between displayed energy and treatment effects and/or decrease of laser efficacy |  Fiber damaged Blast shield mirror broken  Damage of optical components on beam path after photodiode (phd) assembly  Photodiode miscalibration  Misalignment of beam delivery |  Change fiber/ Blast shield Check optic output power with external power meter and phd calibration  Check beam delivery alignment  Check optical components conditions  Substitute damaged parts
Serial port communication fault |  This warning appears when one of the serial communication fails. |  Restart the system  Check the signal connections  Substitute damaged parts
Capacitor voltage error |  Power supply error or broken  Charger error or broken  Wrong electrical connections |  Check line voltage at power supply, charger and discharger input  Check power supply output voltage  Check signal connections  Check high voltage  Substitute damages parts  Check electrical connections
Charger overload |  Charger error or broken  Line voltage too low |  Check line voltage at power supply, charger and discharger input  Check power supply output voltage  Check electrical connections  Check high voltage  Substitute damages parts
High temperature discharger #(n) |  Problem in power electronics  Main Power Supply overtemperature error  Wrong voltage on the capacitors |  Check operating temperature of the laser system is < 30°C  Check fans are working properly  Check high voltage  Substitute damages parts
Energy Low - Lamp #(n) |  Fluctuations of the energy  Power supply error or damaged  Lamp broken or old  Optical damage or misalignment  Photodiode miscalibration or damage |  Wait few seconds until the energy is stabilized Check high voltage  Check optical components  Check photodiode calibration  Check if lamp is ignited  Check flash lamp  Substitute damages parts
Energy High - Lamp #(n) |  Fluctuations of the energy  Power supply error or damaged  Lamp broken or old  Optical damage or misalignment  Photodiode miscalibration or damage |  Wait few seconds until the energy is stabilized Check high voltage  Check optical components  Check photodiode calibration  Check if lamp is ignited  Check flash lamp  Substitute damages parts
Low Energy warning (<20%) - Lamp #(n) |  Fluctuations of the energy  Power supply error or damaged  Lamp broken or old  Optical damage or misalignment  Photodiode miscalibration or damage |  Wait few seconds until the energy is stabilized Check high voltage  Check optical components  Check photodiode calibration  Check if lamp is ignited  Check flash lamp  Substitute damages parts
High Energy warning (>20%) - Lamp #(n) |  Fluctuations of the energy  Power supply error or damaged  Lamp broken or old  Optical damage or misalignment  Photodiode miscalibration or damage |  Wait few seconds until the energy is stabilized Check high voltage  Check optical components  Check photodiode calibration  Check if lamp is ignited  Check flash lamp  Substitute damages parts
Lamp #(n) not ignited |  Discharger error or broken  Power supply error or damaged  Lamp broken or old  Simmering transformer damaged |  Check line voltage at power supply, charger and discharger input  Check power supply output voltage  Check electrical connections  Check high voltage  Check presence of simmering trigger  Check flash lamp  Check lamp simmering and if lamp is not emitted  Substitute damages parts
Simmer #(n) error (Lamp simmering ON when it should be OFF) |  Discharger error or broken |  Check line voltage at discharger input  Check electric signals/connections  Substitute damages parts
Low water flow #(n) |  Low water level  Water tube is twisted  Filter blocked |  Refill water  Check water pipes and filter  Check flowmeters  Check hydraulic circuit condition
Flow-switch #(n) error |  Water flow out of range  Flow meter broken |  Check water pipes  Substitute damages parts
Chiller error (level, temperature) |  Water missing  Environment temperature high  Gas circuit issue |  Check the water level and refill if necessary  Check cooling air circulation and assure the correct distance of the rear panel from the wall
Galvanometer motor error |  Galvo motor broken  Galvo driver broken |  Check connections  Check input board voltage  Check electric signals  Substitute damages parts
Shutter Error |  Incorrect shutter position: Shutter off fault The command sent by the microprocessor was “Shutter close”, but the shutter remains “open”. Shutter not open The command sent by the microprocessor was “Shutter open”, but the shutter remains “close”.  Mechanical interference between actuator and board  Over heated actuator produce its stuck  Damaged board QS05-2H, or QSVM1602 or QSMK1701  Not proper interconnections between microprocessor and shutter |  Check shutter opto-mechanical assembly and its free movement (with machine switched off)  Check electric signals  Substitute damaged parts
The remote interlock contacts are open |  The ext. Interlock is not connected or is poorly connected  The external door safety switch is open |  Check ext. Interlock connection  Check internal connections
Pedal not connected |  Wrong pedal connections |  Check connections
Pedal pressed in STANDBY |  Pedal damaged  Pedal(s) should be pressed at the right moment |  Check connections  Enter in Ready mode and press the footswitch  Substitute damages parts
Blast shield not connected |  Wrong blast shield installation  Blast shield microswitch damaged |  Check Blast shield  Check connections  Substitute damages parts
Fiber not present/connected |  Wrong fiber installation |  Check the fiber connection
Fiber not identified by RFID |  Wrong fiber installation Wrong fiber type  RFID antenna |  Check the fiber connection  Check the RFID frequency  Change fiber
RFID warning |  RFID antenna  Board not working  Wrong fiber type  Wrong fiber code  The fiber is expired |  Check the fiber connection  Check the RFID frequency  Check electrical connections  Check or change fiber  Substitute damaged parts
FW error |  Message after new FW release installation FW issues |  Shut down and restart Call service
SW error |  SW issues |  Shut down and restart Call service

## CYM
PROBLEMA | CAUSA | SOLUZIONE
Key switch does nothing |  Power line is not connected Main line breaker is off Emergency red push button is pressed  Main line input/low voltages connections are not proper  Low voltages Power supply (PS) is damaged  Microprocessor or embedded PC boards are damaged |  Check main line breaker (room and machine)  Defeat the emergency push button  Check main line filter I/O voltage  Check key and emergency push button I/O voltage  Check low voltage PS I/O voltages  Check low voltages on microprocessor board and embedded PC  Substitute damaged parts
No correspondence between displayed energy and treatment effects and/or decrease of laser efficacy |  Fiber damaged Blast shield mirror broken  Damage of optical components on beam path after photodiode (phd) assembly  Photodiode miscalibration  Misalignment of beam delivery |  Change fiber/ Blast shield Check optic output power with external power meter and phd calibration  Check beam delivery alignment  Check optical components conditions  Substitute damaged parts
Serial port communication fault |  This warning appears when one of the serial communication fails. |  Restart the system  Check the signal connections  Substitute damaged parts
Capacitor voltage error |  Power supply error or broken  Charger error or broken  Wrong electrical connections |  Check line voltage at power supply, charger and discharger input  Check power supply output voltage  Check signal connections  Check high voltage  Substitute damages parts  Check electrical connections
Charger overload |  Charger error or broken  Line voltage too low |  Check line voltage at power supply, charger and discharger input  Check power supply output voltage  Check electrical connections  Check high voltage  Substitute damages parts
High temperature discharger #(n) |  Problem in power electronics  Main Power Supply overtemperature error  Wrong voltage on the capacitors |  Check operating temperature of the laser system is < 30°C  Check fans are working properly  Check high voltage  Substitute damages parts
Energy Low - Lamp #(n) |  Fluctuations of the energy  Power supply error or damaged  Lamp broken or old  Optical damage or misalignment  Photodiode miscalibration or damage |  Wait few seconds until the energy is stabilized Check high voltage  Check optical components  Check photodiode calibration  Check if lamp is ignited  Check flash lamp  Substitute damages parts
Energy High - Lamp #(n) |  Fluctuations of the energy  Power supply error or damaged  Lamp broken or old  Optical damage or misalignment  Photodiode miscalibration or damage |  Wait few seconds until the energy is stabilized Check high voltage  Check optical components  Check photodiode calibration  Check if lamp is ignited  Check flash lamp  Substitute damages parts
Low Energy warning (<20%) - Lamp #(n) |  Fluctuations of the energy  Power supply error or damaged  Lamp broken or old  Optical damage or misalignment  Photodiode miscalibration or damage |  Wait few seconds until the energy is stabilized Check high voltage  Check optical components  Check photodiode calibration  Check if lamp is ignited  Check flash lamp  Substitute damages parts
High Energy warning (>20%) - Lamp #(n) |  Fluctuations of the energy  Power supply error or damaged  Lamp broken or old  Optical damage or misalignment  Photodiode miscalibration or damage |  Wait few seconds until the energy is stabilized Check high voltage  Check optical components  Check photodiode calibration  Check if lamp is ignited  Check flash lamp  Substitute damages parts
Lamp #(n) not ignited |  Discharger error or broken  Power supply error or damaged  Lamp broken or old  Simmering transformer damaged |  Check line voltage at power supply, charger and discharger input  Check power supply output voltage  Check electrical connections  Check high voltage  Check presence of simmering trigger  Check flash lamp  Check lamp simmering and if lamp is not emitted  Substitute damages parts
Simmer #(n) error (Lamp simmering ON when it should be OFF) |  Discharger error or broken |  Check line voltage at discharger input  Check electric signals/connections  Substitute damages parts
Low water flow #(n) |  Low water level  Water tube is twisted  Filter blocked |  Refill water  Check water pipes and filter  Check flowmeters  Check hydraulic circuit condition
Flow-switch #(n) error |  Water flow out of range  Flow meter broken |  Check water pipes  Substitute damages parts
Chiller error (level, temperature) |  Water missing  Environment temperature high  Gas circuit issue |  Check the water level and refill if necessary  Check cooling air circulation and assure the correct distance of the rear panel from the wall
Galvanometer motor error |  Galvo motor broken  Galvo driver broken |  Check connections  Check input board voltage  Check electric signals  Substitute damages parts
Shutter Error |  Incorrect shutter position: Shutter off fault The command sent by the microprocessor was “Shutter close”, but the shutter remains “open”. Shutter not open The command sent by the microprocessor was “Shutter open”, but the shutter remains “close”.  Mechanical interference between actuator and board  Over heated actuator produce its stuck  Damaged board QS05-2H, or QSVM1602 or QSMK1701  Not proper interconnections between microprocessor and shutter |  Check shutter opto-mechanical assembly and its free movement (with machine switched off)  Check electric signals  Substitute damaged parts
The remote interlock contacts are open |  The ext. Interlock is not connected or is poorly connected  The external door safety switch is open |  Check ext. Interlock connection  Check internal connections
Pedal not connected |  Wrong pedal connections |  Check connections
Pedal pressed in STANDBY |  Pedal damaged  Pedal(s) should be pressed at the right moment |  Check connections  Enter in Ready mode and press the footswitch  Substitute damages parts
Blast shield not connected |  Wrong blast shield installation  Blast shield microswitch damaged |  Check Blast shield  Check connections  Substitute damages parts
Fiber not present/connected |  Wrong fiber installation |  Check the fiber connection
Fiber not identified by RFID |  Wrong fiber installation Wrong fiber type  RFID antenna |  Check the fiber connection  Check the RFID frequency  Change fiber
RFID warning |  RFID antenna  Board not working  Wrong fiber type  Wrong fiber code  The fiber is expired |  Check the fiber connection  Check the RFID frequency  Check electrical connections  Check or change fiber  Substitute damaged parts
FW error |  Message after new FW release installation FW issues |  Shut down and restart Call service
SW error |  SW issues |  Shut down and restart Call service

## LHV
PROBLEMA | CAUSA | SOLUZIONE
The device cannot start up |  Power line not connected  Main line breaker is off  Emergency push button is pressed down  MPS broken |  Check line cord breaker  Pull up the emergency push button  Check fuses  Substitute MPS
Footswitch Error |  Footswitch disconnected  Defective Footswitch |  Connect the footswitch  Check if footswitch must be replaced
Water Flow Error |  Not enough water  Water filter occluded  Water tube twisted  Water leakage |  Fill water reservoir  Check water filter  Check water tubes  Look for any water leakage
Thermo switch error |  Room temperature too high or air inlet/outlet not free  Temperature sensing disconnected or damaged |  Check cooling circulation and fans  Assure that a correct distance from the rear panel to the wall exists  Check water pump  Check NTC and thermoswitch connection/status.
Flashlamp error |  Flashlamp is exhausted or broken  Power supply |  Check power supply  Substitute the flashlamp
Simmering Error |  Simmering breakage  Flashlamp broken  Trigger transformer failed |  Check the compact power supply H.V.P.S.  Check the flashlamp  Check trigger transformer
OverVoltage, Temp Error |  Problem in power electronics |  Check the compact power supply H.V.P.S.
External Interlock Error |  Interlock not well connected or not connected at all |  Check the connection of the Interlock
Shutter Error |  Incorrect shutter position |  Change the shutter or its electronic board
Energy Warning: high or low energy |  Fluctuation of the energy  Incorrect energy |  Check the photodiode reading  Check the optics and laser head/alignment
No correspondence between displayed energy and application effects |  Damage of the fiber |  Substitute the fiber

## TFL
PROBLEMA | CAUSA | SOLUZIONE
Pressing the ON/OFF button is not producing any detectable effect |  Plug not connected  Main switch not turned on  ON/OFF key not turned in ON position  Emergency LASER STOP activated |  Check for main input  Deactivate Emergency LASER STOP switch  Turn the key on the ON position
Energy error | Fluctuation of energy parameter during laser radiation |  Wait a few seconds for stabilization.  Measure the output power  Check the fiber laser power
Serial port communication fault | • This warning appears when the serial communication from PC to microprocessor fails |  Restart the system  Check the connections
Capacitor voltage error | • Power supply error or broken • Relay board failure • Damaged capacitors |  Check the power supply  Check the power connections  Check the relays board
No correspondence between displayed energy and treatment effects and/or decrease of laser efficacy | •Fiber damaged • Blast shield optic broken | • Check fiber • Check the blast shield
Shutter not closed | • Incorrect shutter position | • Check the shutter and its connection
Shutter not open | • Incorrect shutter position | • Check the shutter and its connection
The remote interlock contacts are open | • The ext. Interlock is not connected or is poorly connected | • Control ext. Interlock connections
Pedal not connected | • Wrong pedal connections | • Check connections
Pedal pressed in STANDBY | • Pedal(s) should be pressed at the right moment | • Enter in Ready mode and press the footswitch
Blast shield wrong installation | • Wrong blast shield installation | • Check Blast shield
Fiber not present/connected | • Wrong fiber installation | • Check the fiber connection
Fiber not identified by RFID | • Wrong fiber installation • Wrong fiber type | • Check the fiber connection • Change fiber
RFID warning | • Wrong fiber type • Change fiber • Wrong fiber type • Wrong fiber code | • Change fiber
FW error | • Message after new FW release installation • FW issues | • Shut down and restart
Error REV24 | No connection 5V to embedded | check the connection of the cables Check the emergency button
Unwanted Emission |  | Check Diode cable on the board
Source Error | No laser output | Restart the system Replace the Laser Source
Very low Energy | Low energy output | Cange fiber Back Reflections Photodiode wrong calibrated

## TFN
PROBLEMA | CAUSA | SOLUZIONE
Pressing the ON/OFF button is not producing any detectable effect | • Plug not connected • Main switch not turned on • ON/OFF key not turned in ON position • Emergency LASER STOP activated | • Check for main input • Deactivate Emergency LASER STOP switch • Turn the key on the ON position
Energy error | Fluctuation of energy parameter during laser radiation | • Wait a few seconds for stabilization. • Measure the output power • Check the fiber laser power
Serial port communication fault | • This warning appears when the serial communication from PC to microprocessor fails | • Restart the system • Check the connections
No correspondence between displayed energy and treatment effects and/or decrease of laser efficacy | • Fiber damaged • Blast shield optic broken | • Check fiber • Check the blast shield
Shutter not closed | • Incorrect shutter position | • Check the shutter and its connection
Shutter not open | • Incorrect shutter position | • Check the shutter and its connection • Control ext. Interlock connections
The remote interlock contacts are open | • The ext. Interlock is not connected or is poorly connected | • Check connections
Pedal not connected | • Wrong pedal connections | • Enter in Ready mode and press the footswitch
Pedal pressed in STANDBY | • Pedal(s) should be pressed at the right moment | • Check Blast shield
Blast shield not present | • Wrong blast shield installation or not installed | • Check the fiber connection • Check the fiber presece sensor
Fiber not present/connected | • Wrong fiber installation • Malfunction of fiber presece sensor • Wrong fiber installation • Wrong fiber type | • Check the fiber connection • Check the fiber presece sensor
Fiber not identified by RFID | • Malfunction of RFID | • Check the fiber connection • Change fiber • Check the RFID
Fiber not usable | • Wrong fiber type • Wrong fiber code • The fiber is expired | • Change fiber • Check fiber
FW CRC32 error | • Message after new FW release installation • FW issues | • Clear the CRC32 error in Service menu
FW error | • FW issues | • Shut down and restart
Error REV24 | No connection 5V to embedded | check the connection of the cables Check the emergency button
Unwanted Emission |  | Check Diode cable on the board
Source Error | No laser output | Restart the system Replace the Laser Source
Very low Energy | Low energy output | Cange fiber Back Reflections Photodiode wrong calibrated

## DNL
PROBLEMA | CAUSA | SOLUZIONE
Pressing the ON/OFF button is not producing any detectable effect | • Plug not connected • Main switch not turned on • ON/OFF key not turned in ON position • Emergency LASER STOP activated | • Check for main input • Deactivate Emergency LASER STOP switch • Turn the key on the ON position
Energy error | Fluctuation of energy parameter during laser radiation | • Wait a few seconds for stabilization. • Measure the output power • Check the fiber laser power
Serial port communication fault | • This warning appears when the serial communication from PC to microprocessor fails | • Restart the system • Check the connections
No correspondence between displayed energy and treatment effects and/or decrease of laser efficacy | • Fiber damaged • Blast shield optic broken | • Check fiber • Check the blast shield
Shutter not closed | • Incorrect shutter position | • Check the shutter and its connection
Shutter not open | • Incorrect shutter position | • Check the shutter and its connection • Control ext. Interlock connections
The remote interlock contacts are open | • The ext. Interlock is not connected or is poorly connected | • Check the shutter and its connection • Control ext. Interlock connections
Pedal not connected | • Wrong pedal connections | • Check connections
Pedal pressed in STANDBY | • Pedal(s) should be pressed at the right moment | • Enter in Ready mode and press the footswitch
Blast shield not present | • Wrong blast shield installation or not installed | • Check Blast shield
Fiber not present/connected | • Wrong fiber installation • Malfunction of fiber presece sensor | • Check the fiber connection • Check the fiber presece sensor
Fiber not identified by RFID | • Wrong fiber installation • Wrong fiber type | • Check the fiber connection • Change fiber • Check the RFID